## Slide 1: The Proton Computed Tomography Apparatus developed by INFN
M. Bruzzi1,2, C. Civinini2, G. Maccioni3, S. Pallotta2,4,5, F. Paulis3, N. Randazzo6, M. Scaringella7, V. Sipala3,8, C. Talamonti2,4,5, E. Vanzi9
1Physics and Astronomy Department, University of Florence, Florence, Italy 2INFN - Florence Division, Florence, Italy
3INFN Cagliari Division, Cagliari, Italy 4Department of Biomedical, Experimental and Clinical Sciences, University of Florence, Florence, Italy
5 SOD Fisica Medica, Azienda Ospedaliero-Universitaria Careggi, Firenze, Italy 6INFN - Catania Division, Catania, Italy
7Information Engineering Department, University of Florence, Florence, Italy 8Chemistry and Pharmacy Department, University of Sassari, Sassari, Italy 9 Fisica Sanitaria, Azienda Ospedaliero-Universitaria Senese, Siena, Italy
Advanced Semiconductor Detectors for Medical Applications
Garching 13/02/2015

## Slide 2: Introduction
Two pCT systems: 5x5cm2; 5x20cm2
Results from the small area system
  Filtered Back Projection
  Algebraic Reconstruction Technique
Large area apparatus configuration
  Tracker (some considerations on MLP)
  Calorimeter
  Present status

## Slide 3: pCT apparatus
PARAMETER | VALUE
Proton beam kinetic energy | ~200 MeV
Proton beam rate | 1 MHz
Spatial resolution | < 1 mm
Electronic density resolution | <1%
Detector radiation hardness | >1000 Gy
Dose per scan | < 5 cGy
P1
P2
P3
P4
z
x
y
Single particle proton tracking: silicon strip detectors → MLP
Residual energy measurement: crystal calorimeter → energy loss
A set of single event information can be processed by appropriate reconstruction algorithms (ART, FBP) to produce tomographic images.

## Slide 4: PRIMA collaboration: small area pCT apparatus
Four x-y silicon microstrip based
tracking planes
Yag:Ce calorimeter
Proton entry and exit positions and directions
Proton residual energy
First test at INFN-LNS:
May 2011
CATANA beam line:
62 MeV protons
used to treat ocular tumors

## Slide 5: Tracker module
Parallel strip read-out
Local data storing during measurement
Ethernet data download at measurement completion
Front-end board
Digital board

## Slide 6: Si Sensor and Front-end ASIC
6.6 x 1.6 mm2
32 inputs - 32 outputs
670 mW power consumption
Vcc=+3.3 V
p on n
single sided
<100>
200mm thick
200mm strip pitch

## Slide 7: Tracker performance
62 MeV proton
minimum released charge
in 200mm of silicon
180 MeV proton
minimum released charge
in 200mm of silicon
Calibration
Efficiency
Signal duration
Noise occupancy per sensor
Data INFN-LNS +TSL (Uppsala)

## Slide 8: Tomografic set-up

## Slide 9: 62 MeV pCT image - FBP
PMMA phantom
36 projection steps: 0°  360°
An average of 950000 events per projection
E0=62MeV INFN-LNS
Filtered Back Projection algorithm
Tomographic equation
(Wang, Med.Phys. 37(8), 2010: 4138)
«projection»
Unknown stopping power distribution (at E0)
Evaluation of the “projection” term (through numerical integration starting from NIST tables and using the measured Eres)
Eres
Wang projection

## Slide 10: 62 MeV pCT image - FBP
No selection cuts has been applied to the data sample
E. Vanzi et al. The PRIMA collaboration: preliminary results in FBP reconstruction of pCT data –RESMDD12 Conference doi:10.1016/j.nima.2013.05.193
Chart: Resolution (no cut)
| 2 | 4 | 6 | 8 | 10 | 12
20/128 | 0.78762 | 0.88283 | 0.93064 | 0.96498 | 0.99158 | 1.0128
Butterworth filter: order 2,
cut-off 20/128 of the Nyquist freq.
noise: 6.3%
noise: 2.4%

## Slide 11: Algebraic Reconstruction Techniques
Iterative algorithm to reconstruct tomographic images from projections
Iterative formula:
xk image vector at iteration k (relative stopping power)
ai ith track length in each pixel  Tracker
bi stopping power integral  Calorimeter
lk relaxing factor (constant value or 0 as ~k-1)
x0 initial image: {0} or approx.
Gordon, R; Bender, R; Herman, GT J. Theor. Biol. (1970) 29 (3): 471–81.
xk+1= xk + lk {bi - <ai,xk> } ai / ǁaiǁ2

## Slide 12: ART images from 62 MeV data
cm
x100mm
ART reconstruction:
14 iteration starting from {0}
~ 106 events per angle
(36 angles)
4mm vertical slice selected
(2D only)
~4’ per iteration CPU time
No MLP is used  to be done
x100mm
PMMA
phantom
1 cm

## Slide 13: ART images from 62 MeV data
Starting image = {0}, still room to improve
Resolution [mm]

## Slide 14: FBP used as seed for ART
FBP initial image 
ART after 14 iterations starting
from FBP seed
x100mm
x100mm
x100mm
x100mm

## Slide 15: ART Resolutions
External edges
Inner holes:
resolution affected by
multiple scattering
Resolution [mm]
Small hole
Large hole
Two different edge
Positions

## Slide 16: pCT upgrade (5x20cm2)
A system similar to the one already tested
  Microstrip tracker
  YAG:Ce calorimeter
But with a 50 x 200 mm2 field of view
On-line data aquisition
1 MHz capability
Rectangular aspect ratio to perform tomographies in slices
Beam pipe
Tracker planes
Phantom
Calorimeter
Silicon sensors

## Slide 17: Tracker plane configuration
2x4 silicon microstrip sensors glued onto a 40x45 cm2 12 layer PCB (components on both sides: front x-coordinate, rear y-coordinate)
  1 PCB per plane instead of 4 as in the previous prototype
6 read-out groups, 256 channels each
One Xilinx Spartan 6 per group for: sampling, event assembly, zero suppression and data transmission
One Spartan 6 for event building, data receiving and buffer, high speed link (8 serial links running at 200Mbs)
Silicon sensors overlapped to avoid blind area (see comments on MLP error envelope toward the end of this talk)
Up to 5 planes could be read-out by a Virtex6 acquisition board

## Slide 18: Fully assembled Tracker plane
Master FPGA
Virtex6
Silicon mstrip
Slave FPGAs
Chip front-end

## Slide 19: Silicon microstrip detectors
Chart: Bias Voltage (V)
| 5 | 10 | 15 | 20 | 25 | 30 | 35 | 40 | 45 | 50 | 55 | 60 | 65 | 70 | 75 | 80 | 85 | 90 | 95 | 100 | 105 | 110 | 115 | 120 | 125 | 130 | 135 | 140 | 145 | 150 | 155 | 160 | 165 | 170 | 175 | 180 | 185 | 190 | 195 | 200
 | 16.679 | 22.76 | 28.8 | 34.8 | 38.78 | 41.4 | 43.65 | 45.63 | 47.42 | 49.06 | 50.51 | 51.89 | 53.12 | 54.27 | 55.3 | 56.25 | 57.13 | 57.96 | 58.76 | 59.54 | 60.36 | 61.05 | 61.78 | 62.5 | 63.17 | 63.86 | 64.53 | 65.15 | 65.8 | 66.41 | 67.02 | 67.58 | 68.18 | 68.72 | 69.26 | 69.82 | 70.34 | 70.87 | 71.4 | 71.89
 | 19.48 | 27.89 | 35.74 | 43.89 | 48.79 | 51.89 | 54.51 | 56.81 | 58.89 | 60.76 | 62.43 | 64.03 | 65.37 | 66.62 | 67.77 | 68.97 | 69.74 | 70.65 | 71.5 | 72.33 | 73.4 | 73.94 | 74.72 | 75.47 | 76.17 | 77.1 | 77.54 | 78.21 | 78.88 | 79.51 | 80.17 | 80.78 | 81.41 | 81.99 | 82.57 | 83.14 | 83.7 | 84.25 | 84.81 | 85.32
 | 15.89 | 21.96 | 27.65 | 33.58 | 37.29 | 39.75 | 41.79 | 43.57 | 45.21 | 46.66 | 48 | 49.25 | 50.36 | 51.39 | 52.36 | 53.25 | 54.07 | 54.86 | 55.64 | 56.36 | 57.13 | 57.82 | 58.51 | 59.19 | 59.81 | 60.45 | 61.07 | 61.69 | 62.31 | 62.89 | 63.47 | 64.02 | 64.58 | 65.14 | 65.64 | 66.18 | 66.66 | 67.19 | 67.7 | 68.21
 | 20.94 | 30.21 | 38.5 | 47.42 | 52.67 | 55.92 | 58.67 | 61.15 | 63.36 | 65.3 | 67.1 | 68.7 | 70.12 | 71.43 | 72.57 | 73.6 | 74.58 | 75.44 | 76.3 | 77.1 | 77.95 | 78.66 | 79.42 | 80.15 | 80.86 | 81.57 | 82.22 | 82.89 | 83.58 | 84.18 | 84.86 | 85.44 | 86.05 | 86.68 | 87.24 | 87.83 | 88.38 | 88.94 | 89.5 | 90.02
 | 16.047 | 22.3 | 28.18 | 34.3 | 38.27 | 40.74 | 42.91 | 44.9 | 46.61 | 48.14 | 49.81 | 50.88 | 52.1 | 53.16 | 54.21 | 55.12 | 55.99 | 56.82 | 57.64 | 58.38 | 59.21 | 59.9 | 60.61 | 61.33 | 61.95 | 62.62 | 63.25 | 63.88 | 64.52 | 65.11 | 65.7 | 66.26 | 66.87 | 67.4 | 67.93 | 68.5 | 69.02 | 69.53 | 70.07 | 70.56
 | 27.65 | 41.81 | 54.42 | 67.94 | 74.86 | 79.29 | 82.83 | 86.04 | 88.79 | 91.3 | 93.52 | 95.49 | 97.25 | 98.79 | 100.25 | 101.4 | 102.49 | 103.51 | 104.47 | 105.4 | 106.38 | 107.23 | 108.08 | 108.9 | 109.67 | 110.43 | 111.19 | 111.95 | 112.71 | 113.4 | 114.14 | 114.8 | 115.48 | 116.17 | 116.8 | 117.48 | 118.1 | 118.76 | 119.4 | 119.99
 | 21.09 | 30.7 | 39.32 | 48.63 | 53.67 | 57.18 | 60.02 | 62.52 | 64.76 | 66.74 | 68.54 | 70.16 | 71.63 | 72.94 | 74.11 | 75.16 | 76.09 | 77 | 77.87 | 79.01 | 79.58 | 80.51 | 81.11 | 81.86 | 82.56 | 83.29 | 83.97 | 84.69 | 85.37 | 86 | 86.65 | 87.24 | 87.87 | 88.47 | 89.05 | 89.65 | 90.21 | 90.9 | 91.4 | 91.93
 | 16 | 21.98 | 27.42 | 33.3 | 37.22 | 39.71 | 41.8 | 43.65 | 45.33 | 46.86 | 48.26 | 49.52 | 50.69 | 51.77 | 52.76 | 53.68 | 54.55 | 55.36 | 56.14 | 56.9 | 57.71 | 58.4 | 59.11 | 59.81 | 60.47 | 61.14 | 61.77 | 62.42 | 63.06 | 63.64 | 64.24 | 64.8 | 65.38 | 65.94 | 66.47 | 67.02 | 67.53 | 68.08 | 68.61 | 69.11
 | 16.895 | 23.53 | 30.16 | 36.92 | 41.42 | 44.28 | 46.72 | 48.82 | 50.74 | 52.46 | 54.06 | 55.49 | 56.81 | 58.07 | 59.16 | 60.2 | 61.13 | 62.06 | 62.92 | 63.75 | 64.64 | 65.35 | 66.12 | 66.86 | 67.54 | 68.26 | 68.89 | 69.59 | 70.25 | 70.87 | 71.52 | 72.12 | 72.74 | 73.35 | 73.91 | 74.51 | 75.05 | 75.63 | 76.22 | 76.75
 | 15.534 | 21 | 26.63 | 32.68 | 36.26 | 38.71 | 40.77 | 42.58 | 44.24 | 45.72 | 47.07 | 48.34 | 49.49 | 50.54 | 51.51 | 52.44 | 53.3 | 54.09 | 54.89 | 55.68 | 56.47 | 57.19 | 57.88 | 58.56 | 59.21 | 59.88 | 60.51 | 61.15 | 61.77 | 62.35 | 62.93 | 63.48 | 64.07 | 64.63 | 65.17 | 65.73 | 66.22 | 66.75 | 67.27 | 67.75
 | 16.125 | 21.7 | 27.2 | 32.99 | 36.91 | 39.42 | 41.8 | 43.42 | 45.11 | 46.62 | 48.03 | 49.3 | 50.47 | 51.57 | 52.57 | 53.48 | 54.35 | 55.19 | 55.96 | 56.73 | 57.55 | 58.24 | 58.95 | 59.66 | 60.3 | 60.97 | 61.59 | 62.23 | 62.87 | 63.47 | 64.06 | 64.6 | 65.17 | 65.74 | 66.26 | 66.79 | 67.32 | 67.9 | 68.38 | 68.89
 | 18.68 | 26.44 | 34.12 | 42.99 | 46.97 | 50.1 | 52.67 | 55.01 | 57.04 | 58.9 | 60.6 | 62.12 | 63.52 | 65.13 | 65.98 | 67.06 | 68.06 | 69 | 69.89 | 71.02 | 71.7 | 72.46 | 73.25 | 74.03 | 74.74 | 75.48 | 76.17 | 76.87 | 77.55 | 78.21 | 78.88 | 79.51 | 80.15 | 80.77 | 81.39 | 82.01 | 82.59 | 83.19 | 83.77 | 84.32
36 p on n single-sided silicon microstrip detectors (HPK):
  51.2 x 51.2 mm2 active area
  <100> crystal type
  320 mm thickness
  200 mm pitch
  256 channels
  I(Vfd)~2-3nA/cm2
No bad strips for all 36 detectors (9216 strips)
Chart: 
| 75 | 80 | 85 | 90 | 95 | 100
Vfd | 0 | 3 | 28 | 1 | 4 | 0
Full Depletion Voltage (V)

## Slide 20: Front-end ASICs
32 channels parallel output discriminator (evolution of the previuos one)
Channel-by-channel 8-bit thresholds loaded via I2C
AMS–CMOS 0.35mm

## Slide 21: Read-out group
Front-end chips
FPGA slave
Silicon microstrip detectors
Power regulators

## Slide 22: Tracker DAQ architecture
Central acquisition
Virtex 6
ML605 board
Calorimeter DAQ
Trigger
enable
Trigger
7
gen
SS 1
SS 6
.
.
.
256
256
SM
serial link data/clk
200 Mbs
SSD 1
SSD 6
256
256
Front end 1
Front end 6
tracker plane
SS 1
SS 6
.
.
.
256
256
SM
serial link data/clk
200 Mbs
SSD 1
SSD 6
256
256
Front end 1
Front end 6
tracker plane
SS 1
SS 6
.
.
.
256
256
SM
serial link data/clk
200 Mbs
SSD 1
SSD 6
256
256
Front end 1
Front end 6
tracker plane
SS 1
SS 6
.
.
.
256
256
SM
6 readout links
200(x8) Mbs each
SSD 1
SSD 6
256
256
Front end 1
Front end 6
tracker plane
4/5 readout links 1.6 Gbs each
The front end of each detector will read-out by a Xilinx Spartan 6 FPGA (SS)
When a trigger occurs the read-out unit containing at least one hit will send the data to a central FPGA (SM)
The SM will asynchronously send the data to the central acquisition board
Trigger
gen
Trigger enable
7
1Gbs Ethernet
2GB DDR3 RAM

## Slide 23: Test results with high intensity 90Sr source
Source profile
Two not working strips:
one because of sensor
one because of front-end chip
(one ~hot)
# strip
# cluster
Number of clusters (electrons) per
event <Nclu>=4.2
2.5ms DAQ window  ~ 2 MHz of e-
Presently @350kHz max. trigger rate

## Slide 24: MLP in a realistic pCT geometry
MLP example with 200MeV
kinetic energy protons in
20cm of water:
Entry: Y(0) = 0.2cm
Y’(0) = -10mrad
Exit: Y(20) = -0.1cm
Y’(20) = +10mrad
Silicon microstrip detectors:
320mm thick
200mm strip pitch
MLP error envelope plus contributions from detector position measurement error (~ pitch/√12) and MCS inside the silicon sensors 
The sensor thickness contribution affects only the MLP error at the edge of the phantom
s ~ 150-250mm
200MeV in
90MeV out
Starting from D.C. Williams Phys. Med. Biol. 49 (2004) and R.W. Shulte at al. Med. Phys. 35 (11) (2008)
5 cm of air have been inserted in front and behind the 20cm H2O phantom

## Slide 25: YAG:Ce calorimeter
2x7 YAG:Ce Crystals Array
Size: 3x3x10cm3
CHASSIS NI PXIe-1071
RT Controller
NI PXI-8102
FlexRIO
NI PXIe-7962R
Ad.Mod.
NI-5751
14 Analog Channels
Tracker
7Dig I/O GEN
Dig. Trigger
Disable Trigger
Silicon Photodiodes 1.8x1.8cm2
Fast Charge Amplifier + Shaper
x14
Data Acquisition System
Parallel read-out
Sampling: 5MS/s
24 Samples x event

## Slide 26: New YAG:Ce calorimeter data
ADC counts
60 MeV
Double protons=120 MeV
Noise = 10 MeV
s=3.9%

## Slide 27: Energy distribution
Preliminary
Uniformity od the calorimeter energy distribution as function of the track impact point as measured by the tracker
E=60MeV
Preliminary crystal calibration

## Slide 28: Medium term plan
Small area system
  Implement MLP into ART (+3D reconstruction)
  Analize Svedberg Laboratory run (Uppsala) at 180 MeV data
  Analize INFN-LNS run (Catania) at 60 MeV with multi-material phantom
  A second run at Uppsala is foreseen this summer
Extended view system
  Validate the tracker plane hardware at 60 MeV
  Production of the 4/5 final tracker planes
  Integration with the calorimeter read-out